Qubit DNA quantification protocol

Standard Operating Procedure #IV.01 v.20250723

AG Sociality & Evolution (Kopp), University of Konstanz

Responsible: 		Dr. Gisela H. Kopp

Date created: 		2022-06-02

Created by: 		Qubit HS Assay kit protocol

Revision #:		5

Revision Date:	 	2025-07-23

Revised by:		GHK

Major points of rev.:	adjust to improved 1X assay

Related protocols:

- NA

# Quantification of DNA using Qubit assay kit

Protocol for quantifying DNA with Qubit high sensitivity (HS) or broad range (BR) assay.

Based on Qubit 1X dsDNA BR assay manual.

## Preconditions and lab environment

- You need to have received lab instructions before starting to work in the lab

  - Z838 (Dr. Gisela H.Kopp, gisela.kopp@uni-konstanz.de)

  - Limnological Institute (Dr. Azade Voolstra, azade.voolstra@uni-konstanz.de)

  - MRC (N.N.)

- Follow Betriebsanweisung_Z838_2022 and Risk_Assessment_biological-samples for working in clean room and with biological samples to ensure decontaminated work space.

- Wear lab coat with cuffed sleeves, sleeve covers, safety glasses, disposable mask and extended gloves throughout the whole procedure. It is critical to avoid contamination.

- Keep a detailed protocol in the hardcopy lab book (needs to stay in the lab) and in the online lab book “labfolder”.

- Note all dates and results in the project overview table.

- Give notice if lab materials run low.

### Time estimate:

- 30 min

## Required instruments and materials

|  | Manufacturer & Article number | Amount total [t]/ per sample [ps] | location |
| --- | --- | --- | --- |
| Equipment |  |  |  |
| Pipettes (volumes 2.5ul, 100ul, 1000ul) | Eppendorf | 1 |  |
| minicentrifuge for 0.5ml tubes |  |  |  |
| Epi rack |  |  |  |
| Qubit instrument |  |  |  |
| Plastic ware |  |  |  |
| 2 ml tubes | Eppendorf |  |  |
| Pipette filter tips 1000ul | Eppendorf/sigmal aldrich | 1 |  |
| Pipette filter tips 2.5ul | Eppendorf/sigmal aldrich | 1 |  |
| Pipette filter tips 100ul | Eppendorf/sigmal aldrich | 1 |  |
| Qubit assay tubes | Thermofisher Scientific Q32856 | 1 |  |
| Reagents |  |  |  |
| Qubit 1X dsDNA HS Assay-Kit (incl. Working Solution, Standard #1, Standard #2) | ThermoFisher Scientific (100) Q33230 (500) Q3231 |  | Working solution RT, standards 2-8°C |
| Qubit 1X dsDNA BR Assay Kit (incl. Working Solution, Standard #1, Standard #2) | ThermoFisher Scientific (100) Q33265 (500) Q33266 |  | Working solution RT, standards 2-8°C |

## Procedure

DNA quant. number: _______________________________		Date:__________________________________________

Personnel: ______________________________________		Sample Information: Number of samples (=n): ________

| Sample ID | Species | Collection time | Collection Date | Collection location | DNA extraction date | Sample volume | DNA extract ID | DNA conc. (ng/ul) | comment |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Blank | x | x | x | x |  |  |  |  |  |

Comments:

### Before starting

- Let standards and samples equilibrate to room temperature

### Preparing standards and samples

- Set up and label Qubit tubes for the number of samples you are quantifying and 2 additional tubes for use as standards.

- Pipett 380ul + n * 199ul (where n is the number of samples to measure) 1X Working solution into a separate tube to avoid contamination of the stock working solution.

[e.g. 2768ul for 12 samples and 2 standards]

- Pipett working solution into Qubit tubes: 190ul for standards and 199ul for samples

- Add 10ul Standard #1 and 10ul Standard #2 to the respective tubes

- Add 1ul sample to the respective designated tube.

*Note: The buffer and dye solution is also light sensitive. Wrap a sheet of foil around it for the duration of the protocol and return it to the drawer immediately after use.

- Vortex all tubes for 5sec

- Spin all tubes in the small centrifuge very briefly to get rid of bubbles. It is very important that all the bubbles are gone prior to running the assay.

*Note: If bubbles result from vortexing samples, gently flick the sides of the tube. Be sure to wait until all bubbles are dissipated before running an assay on the Qubit.

- Incubate all tubes for 2 minutes.

### Read standards and samples

- Plug in Qubit machine. On the home screen, touch DNA and select appropriate assay type. Touch “Read Standards” to proceed.

- Insert tube containing Standard #1 into the sample chamber, close the lid, and touch “Read standard”. When completed, remove Standard #1

- Insert tube containing Standard #21 into the sample chamber, close the lid, and touch “Read standard”. When completed, remove Standard #2

- Touch “Run samples”, select the appropriate sample volume (1ul) and units (ng/ul) on the assay screen.

- Insert one sample tube, close the lid, and touch “Read sample”. Write down concentration. When completed, remove sample. Repeat until all samples have been read.